## Slide 1: System aberration correction in a light sheet microscope for whole brain imaging
Caroline Müllenbroich, Ludovico Silvestri, Leonardo Sacconi & Francesco S. Pavone
LENS, European Laboratory for Nonlinear Spectroscopy, University of Florence, Italy

## Slide 2: 3D imaging of structurally intact clarified brains
Comprehensive mapping and quantification of neuronal projections in the central nervous system
Imaging of large (>1cm3), clarified mouse brains with microscopic resolution
Light-sheet microscopy: single axis scan creates light sheet, projection into specimen at 90° to detection objective
Stack acquisition for 2-3 days, 2-3TB per dataset, 2 weeks stitching
Whole mouse brain [cm]
Pyramidal neuron [mm]
Synapses[µm]

## Slide 3: Tissue clearing
PBST
30% TDE
63% TDE
CLARITY transforms intact tissue into nano-porous, hydrogel-hybridized, lipid-free form
  Step 1: formaldehyde crosslinks the hydrogel monomers to proteins
  Step 2: polymerization of the biomolecule-conjugated monomers into a hydrogel mesh
  Step 3: extraction of the lipids with ETC
  Step 4: refractive-index matching with water-soluble clearing agent 2,2'-thiodiethanol (TDE) (n=1.45 with 63%TDE in PBS).
Chung et al., Nature 2013
Costantini et al., Sci. Rep. accepted

## Slide 4: Sample preparation
PBST
30% TDE
63% TDE
Transgenic animals
Mounting
CLARITY protocol
Refractive index matching
with TDE
Chung et al., Nature 2013
Costantini et al., Sci. Rep. accepted

## Slide 5: Sample chamber
z
x

## Slide 6: Specialised high RI immersion objective
Objective:
  correction collar for n = 1.41 - 1.52
  Olympus, XLSLPLN25XGMP
  25x magnification
  1.0 NA
  EFL 7.2mm
  WD 8mm
  Depth of field: ≈1µm
  Back aperture: 14.4mm
  FN 18mm

## Slide 7
Whole mouse brain tomograhpy
PV
GAD
PI
B
C
PV
D
GAD
E
PI
A
Transgenic mouse brains treated with CLARITY and cleared with TDE 63% imaged with Olympus, 25X objective.
3D rendering of a parvalbumin-dTomato brain.
3D rendering of stacks from PV-dTomato mouse brain, GAD-dTomato mouse brain, PI stained mouse brain, scale bar = 400µm.
(C,D,E) High resolution insert, of the stack corresponding to red boxes in C, scale bar = 100µm.

## Slide 8: FITC-albumin labelling of vasculature
Tomography of one rat hemisphere
Additional perfusion with monomeric solution of FITC-albumin to stain vasculature
Automatic segmentation of vasculature fails with highly aberrated images

## Slide 9: Adaptive Optics
  Substantial optical aberrations arise also in cleared samples
  In light-sheet microscopy illumination and detection light paths are completely uncoupled from each other Each light path encounters its own aberrations
  Adaptive Optics to correct for system- and sample-induced aberrations in the detection path
  Principle: active element which introduces equal but opposite distortion to that introduced by the sample/system
  Active element: deformable membrane mirror
  Wavefront is corrected upon reflection off surface by acting on optical path difference
Zeng et al., Biomed Opt Ex 2012
Notes: Fixed brainbow mouse hippocampus mounted in vectrashield (TPEF imaging), ast, coma, sph, total aberr amplitude

## Slide 10: Deformable mirror
BMC multi | MEMS
Technology | el.static
Coating | Al
Actuators | 140
Stroke | 3.5 μm
Pitch | 400µm
Av. Step size | < 1nm
Actuated area (D) | 4.4mm
AR coating | 400-1100nm
Frame rate (max) | 34kHz

## Slide 11
Vignetting and barrel aberrations
Reduction of beam diameter from 14.4 mm BFP to 4.4mm active area of deformable mirror
Tube lens of 300mm necessitated 80mm lens infront of deformable mirror
FN: 18  field of view of ≈ 720µm
Strong vignetting and barrel aberration at FOV edge

## Slide 12
Plössl-type lenses
Negrean et al., Biomed Opt Ex 2014
Reduction of beam diameter from 14.4 mm BFP to 4.4mm (30%) active area of deformable mirror
Tube lens of 300mm necessitated 80mm lens infront of deformable mirror
FN: 18  field of view of ≈ 720µm
Strong vignetting and barrel aberration at FOV edge

## Slide 13
Plössl-type lenses
Negrean et al., Biomed Opt Ex 2014
Reduction of beam diameter from 14.4 mm BFP to 4.4mm (30%) active area of deformable mirror
Tube lens of 300mm necessitated 80mm lens infront of deformable mirror
FN: 18  field of view of ≈ 720µm
Plössel-type lens assembled from two 100mm achromats with a focal length of 50mm
Strong vignetting and barrel aberration at FOV edge

## Slide 14
v
f 100
f 50
f 100
f 200
f 140
f 200
f 200
f 350
405 nm
445 nm
491 nm
515 nm
561 nm
LM
I1
I2
PBS
f 150
f 150
galvo
galvo
filter
wheel
f 350
LP 552
LP 503
LP 466
LP 427
AOTF
Camera
f 50
f 180
DMM
Illumination
Light-sheet generation
Detection
Olympus
1.0NA
25X
8mm WD
Nikon
10x
0.3NA
17.5mm WD

## Slide 15: Spatial arrangment of the LSFM

## Slide 16: Preliminary trial of DMM capabilities
Before AO
After AO
Gaussian excitation line imaged in Flurescein solution
Gaussian fitting of line profile over entire FOV at run time
Optimised with random search algorithm, Merit factor: Amplitude/STD
Moderate improvement through AO, system is well aligned

## Slide 17
300V
0V
Correction of intentionally induced aberration
Use of correction collar to intentionally introduce spherical aberration
Limited stroke (3.5µm) of the DMM becomes apparent
DMM did not fully compensate for introduced aberrations
Silvestri et al., Microsc Res Tech 2014
Simulation of aberrated PSF assuming perfect objective lens imaging into a slab of refractive index mismatched material.
By correcting up to the second order spherical aberration, a quasi-diffraction-limited regime can be recovered.

## Slide 18
Correction of intentionally induced aberration
Reduce amount of large-amplitude, lower-order aberrations the DMM is required to correct
  place detection objective on translation stage to adjust defocus at runtime
  Adjust correction collar of objective to mid-stack position
Silvestri et al., Microsc Res Tech 2014

## Slide 19: Conclusion
  Even if the brains we image are made transparent, their large size means that also comparatively small heterogenities in the refractive index matching sum up to large optical path differences and therefore strong aberrations at large imaging depth
  Introduction of AO elements can induce aberrations (barrel, vignetting)
AO is capable of improving image quality in our system if we take precautions to avoid running out of stroke

## Slide 20: Outlook
  Towards aberration correction at run time
  Implementation of a wavefront sensor and a strategy for guide star creation
    Beads, nonlinear guide star, small native structures...

## Slide 21: Acknowledgments
Ludovico Silvestri
Irene Costantini
Leonardo Sacconi
Francesco Pavone
Paolo Frasconi
Marco Paciscopi
Giulio Iannello
Alessandro Bria
Leonardo Onofri
Bruno Weber
Matthias Schneider
Jean-Piere Ghobril
Marcel van ‘t Hoff
Funding

## Slide 22: Outline
Whole mouse brain imaging for quantifiable datasets
  Tissue clearing
  Optical aberrations
Microscope Setup
  Implementation of the adaptive optics component
Preliminary trials
Conclusion and outlook

## Slide 23: Edge sharpness
ERF fitting in ROI at run time
Merit factor: Amplitude/STD
USAF test chart
After AO
Before
AO
Boone et al., Med Phys 1994

## Slide 24: Random search algorithm
Maximise merit factor
  Total image intensity
  SNR
  Edge sharpness
  ..
xy

## Slide 25: Random search algorithm
Maximise merit factor
  Total image intensity
  SNR
  Edge sharpness
  ..
xy

## Slide 26: Bleaching compensation
Loop executed within random search algorithm
Periodically re-normalise merit factor (MF)
Circular buffer to store MF (m=5)

## Slide 27: Bleaching compensation
Loop executed within random search algorithm
Periodically re-normalise merit factor (MF)
Circular buffer to store MF (m=5)

## Slide 28
Realtime wavefront sensing in a SPIM microscope, and active aberration tracking (Invited Paper) Paper 9335-9 Time: 1:20 PM - 1:50 PM Author(s): Jonathan M. Taylor, Univ. of Glasgow (United Kingdom); Christopher D. Saunter, Cyril J. Bourgenot, Gordon D. Love, John M. Girkin, Durham Univ. (United Kingdom)

## Slide 29: detection focus adaptation
optimizing parameters at every mm of tissue for all tiles
linearly interpolating the values
optimal parameters identified by finding maxima in a specified neighbourhood corresponding to image quality measure
to automate this process: optical focus-quality measure is the ratio of high frequency and low frequency signal in Fourier space

## Slide 30
Pulse amplifier
PC1
NAS
EOM
Camera
AOTF
RF generator
PC2
DAQ
RS-232 Hub
Galvo 1
Galvo 2
405 nm laser
445 nm laser
491 nm laser
515 nm laser
561 nm laser
1 Gb/s ethernet
10 Gb/s SFP+
CameraLink
Co-axial
HV co-axial
USB
RS-232
RF co-axial
Internal bus
Connections
Supercomputing Centre
RS232 Daisy Chain
Servo motor
Translation x
Servo motor
Translation y
Servo motor
Translation z
Rotation stage
Servo motor

## Slide 31
Camera
CameraLink
board
10Gb network card
System
bus
Workstation
Laboratory
10 Gb Optical fibres
Raw data stream
Data conversion stream
Local SSD RAID 0
10.9TB
Network-attached storage
32.4 TB
Remote repository
CINECA
Transfer to remote repository

## Slide 33
“Large clarified samples can have marked RI inhomogeneity, resulting in the need for correction of misalignment of illumination with the focal plane of the detection objective. We achieve this with a linear adaptive calibration procedure before starting the imaging experiment”
Initial alignment:
Red: focal plane detection objective
Blue:digitaly scanned light sheet
During image acquisition:

## Slide 34: Negrean, Biomed Opt Ex 2014