## Slide 1: System aberration correction for whole brain imaging with a light sheet microscope
M.Caroline Müllenbroich, Ludovico Silvestri, Leonardo Sacconi & Francesco S. Pavone

## Slide 2: Outline
intro
Microscope Setup
  Light-sheet microscopy
  Optical setup
  Sample preparation
Adaptive Optics
  Introduction
  Wavefront sensorless aberration correction
  Random Search Algorithm
Results
  System aberration correction
Outlook

## Slide 3: Light-sheet microscopy
Single axis scan creates light sheet
Projection into specimen at 90° to detection objective
Imaging of large, clarified mouse brains
3D datasets through stitching of adjacent stacks

## Slide 4
v
f 100
f 50
f 100
f 200
f 140
f 200
f 200
f 350
405 nm
445 nm
491 nm
515 nm
561 nm
LM
I1
I2
PBS
f 150
f 150
galvo
galvo
filter
wheel
f 350
LP 552
LP 503
LP 466
LP 427
AOTF
Camera
f 50
f 180
DMM
Illumination
Light-sheet generation
Detection
Olympus
1.0NA
25X
8mm WD
Nikon
10x
0.3NA
17.5mm WD

## Slide 5: Spatial arrangment of the LSFM

## Slide 6: Sample preparation
Transgenic animals
Mounting
CLARITY protocol
Refractive index matching
PBST
30% TDE
63% TDE

## Slide 7: Tissue clearing
PBST
30% TDE
63% TDE
[Chung et al. Nature 2013]
CLARITY transforms intact tissue into nano-porous, hydrogel-hybridized, lipid-free form
fully assembled, optically transparent and macromolecule-permeable
  Step 1: formaldehyde crosslinks the tissue and the hydrogel monomers to proteins
  Step 2: polymerization of the biomolecule- conjugated monomers into a hydrogel mesh
  Step 3: extraction of the lipids with ETC
  Step 4: refractive-index matching with water-soluble clearing agent 2,2'-thiodiethanol (TDE) (n=1.45 with 63%TDE in PBS).

## Slide 8: Sample chamber
z
x

## Slide 9: Olympus high RI immersion objective
correction collar for n = 1.41 - 1.52
Olympus, XLSLPLN25XGMP
25x magnification
1.0 NA
EFL 7.2mm
WD 8mm
depth of field: 1.1µm

## Slide 10
Whole mouse brain tomograhpy
PV
GAD
PI
B
C
PV
D
GAD
E
PI
A
Transgenic mouse brains treated with CLARITY and cleared with TDE 63% imaged with Olympus, 25X objective.
3D rendering of a parvalbumin-dTomato brain.
3D rendering of stacks from PV-dTomato mouse brain, GAD-dTomato mouse brain, PI stained mouse brain, scale bar = 400µm.
(C,D,E) High resolution insert, of the stack corresponding to red boxes in C, scale bar = 100µm.

## Slide 11: FITC-albumin labelling of vasculature
Tomographie of one rat hemisphere
Additional perfusion with monomeric solution of FITC-albumin to stain vasculature
Automatic segmentation of vasculature fails with highly aberrated images

## Slide 12: Aberrations in SPIM
In light sheet microscopy illumination and detection light paths are completely uncoupled from each other
Each light path encounters its own aberrations
Each focal plane is independently affected
Spherical aberration

## Slide 13: Adaptive Optics (AO)
  Compensation of optical aberrations arising also in cleared samples
  Principle: use of shaped wavefront with equal but opposite distortion to that introduced by the sample
  Wavefront generated by deformable membrane mirror (DMM)
  DMM shape determined by random search optimisation algorithm

## Slide 14
Start: initial shape from LUT
Select Random Actuator
Selection weighed with Gaussian distribution choosing central actuators more often than those in periphery
Change Actuator Voltage
Change limited by percentage of total actuator stroke
SOLUTION
Voltage change is accepted
No
Yes
Record Merit Factor
Improvement?
Iterations
Voltage reverts to previous value
At least a 100 successful iterations?
Yes
No

## Slide 15
Continue optimisation
No
Yes
Improvement in
MF within last n iterations?
Activate compensation:
Measure MF for m iterations
Calculate mean and σ of m measurements
Replace best MF with mean
Keep best MF
Apply best mirror shape
Resume optimisation

## Slide 16: Deformable Membrane Mirror
BMC multi | MEMS
Technology | el.static
Coating | Al
Actuators | 140
Stroke | 3.5 μm
Pitch | 400µm
Av. Step size | < 1nm
Actuated area (D) | 4.4mm
AR coating | 400-1100nm
Frame rate (max) | 34kHz

## Slide 19
shape
mid

## Slide 21
Realtime wavefront sensing in a SPIM microscope, and active aberration tracking (Invited Paper) Paper 9335-9 Time: 1:20 PM - 1:50 PM Author(s): Jonathan M. Taylor, Univ. of Glasgow (United Kingdom); Christopher D. Saunter, Cyril J. Bourgenot, Gordon D. Love, John M. Girkin, Durham Univ. (United Kingdom)

## Slide 22: Acknowledgments
Ludovico Silvestri
Irene Costantini
Leonardo Sacconi
Francesco Pavone
Paolo Frasconi
Marco Paciscopi
Giulio Iannello
Alessandro Bria
Leonardo Onofri
Bruno Weber
Matthias Schneider
Jean-Piere Ghobril
Marcel van ‘t Hoff

## Slide 23: Our latest component: detection focus adaptation
optimizing parameters at every mm of tissue for all tiles
linearly interpolating the values
optimal parameters identified by finding maxima in a specified neighbourhood corresponding to image quality measure
to automate this process: optical focus-quality measure is the ratio of high frequency and low frequency signal in Fourier space

## Slide 24
Pulse amplifier
PC1
NAS
EOM
Camera
AOTF
RF generator
PC2
DAQ
RS-232 Hub
Galvo 1
Galvo 2
405 nm laser
445 nm laser
491 nm laser
515 nm laser
561 nm laser
1 Gb/s ethernet
10 Gb/s SFP+
CameraLink
Co-axial
HV co-axial
USB
RS-232
RF co-axial
Internal bus
Connections
Supercomputing Centre
RS232 Daisy Chain
Servo motor
Translation x
Servo motor
Translation y
Servo motor
Translation z
Rotation stage
Servo motor

## Slide 25
Camera
CameraLink
board
10Gb network card
System
bus
Workstation
Laboratory
10 Gb Optical fibres
Raw data stream
Data conversion stream
Local SSD RAID 0
10.9TB
Network-attached storage
32.4 TB
Remote repository
CINECA
Transfer to remote repository

## Slide 28
“Large clarified samples can have marked RI inhomogeneity, resulting in the need for correction of misalignment of illumination with the focal plane of the detection objective. We achieve this with a linear adaptive calibration procedure before starting the imaging experiment”
Initial alignment:
Red: focal plane detection objective
Blue:digitaly scanned light sheet
During image acquisition: